Keithley Repair memory

Start: 12 March 2014

Recently I bought an old Keithley 199 multimeter in an auction at e-bay. It didn't work, but it was cheap. Further analysis showed an error on the digital board. The VIA (peripheral controller) was broken.

I know, I could bough a new via (they are no more produced, I think, but they are still available from resellers on e-bay). But, as an engineer, that solution wasn't satisfactory. I wanted to bring the old multimeter to new standards (at least in the digital side). So with a lot of naivety, I decided to replace the broken board with my own creation.

Actually, the good think with old measurement equipment, is that almost all good manufacturers published pretty detailed service manuals, schematic inclusive. Paraphrasing Feynman, to have the map of the cat is the best form to start.

The Digital interface couldn't be more simple. In order to reduce the number of control lines between digital and analog boards, the Keithley 199 uses a serial interface. Pretty basic, nothing to fancy like spi or i2c. Three control lines go from digital to analog board and one goes back into the digital world. The firs three are serial data which are clocked into shift registers in the analog board, which in turn, control the different switch for range selection, ohms measurements, and analog digital conversion. The line which goes back provides the digital value of the analog signal, the output of the analog digital converter.

Well, this in theory. Now it is time to go into the finer details. One thing, which I personally find fascinating, is that the analog digital converter is actually implemented with discrete parts, and actually controlled from the processor. Nowadays this is no more done, as there is frequently a custom ASIC or an FPGA who does the control of the integrating analog digital converter.

The analog digital converter is from the charge-balancing variety. In order to explain how it works, may be is easier to start with a close cousin, the double ramp integrating analog to digital converter.

The double ramp integrating converter, as it name hints, works integrating the unknown voltage for a fixed amount of time, and then discharging the integration capacitor connecting a reference voltage of opposite polarity and measuring the time the output of the integrator takes to reach zero. This process is shown graphically on figure x. It shows a basic integrator circuit where the voltage output ramps up as function of the unknow voltage and time. The relationship between unknown voltage and time is given by following formula:

In the ramp down part, an know reference voltage, of reverse polarity as the input signal, is connected to the input. The output voltaje ramps down until it reach zero. The output voltage during this ramp down period is given by following formula:

Relating equations 1 and 2, we can calculate the input voltage as function of the time passed between start of the ramp up period and end of the ramp down period:

Easy, isn't it. At least in theory. But as always the devil is in the details. We will talk about this in a moment.

The double ramp converter is easy to implement, and as we can see in formula x, it is almost independent of the actual phisical components. And I say almost, because even thought the values of the components are left out of the equation, this equation assumes ideal integrator behavior in first place. And, oh surprise, it is not!.

The first problem is related to the limited supply voltage. The integrator is, of course, an electronic circuit, and it needs to get its supply voltage from somewhere. If we let connected the unknown voltage at the input of the integrator (see figure x), the output voltage will ramp-up but only until the output voltage reaches nearly the supply voltage, a situation know as saturation.

The second problem comes from the integrating capacitor. I know, it comes a little bit of a surprise, because there is no capacitor in the equation for the output voltage (eq. y), but remember from which equations it comes from:

1 y 2

Did you see?. When you equal these two equations in order to get the output voltage, you make the assumption that the capacitor in the interval one is the same that is on interval two, and it is not always the case. Due to a phenomenon know as “Dielectric absorption”, the capacitor will change its behavior with time.

Here comes the charge balancing in the game. Charge balancing works similar as double slope, first an unknown voltage is integrated, but instead of let the output go indefinite up, a circuit compares the output and when it reaches a defined level, and then, a reference voltage is applied in the input which causes the output to ramp down. When the output voltage has ramp down for a little while, we disconnect again the reference voltage, which causes the output to ramp up again, as the unknown voltage is still present in the input. If we track carefully how much time has the output voltage ramped down, we can also calculate the unknown input voltage.